# Candour

Product & Compliance Specification

| Field | Value |
| --- | --- |
| Document title | Candour — Product & Compliance Specification |
| Document ID | CND-SPEC-001 |
| Version | v0.2 (supersedes v0.1, “ReArm Snap’n’Track” / RST-SPEC-001) |
| Classification | Internal — Pre-Certification Draft |
| Prepared by | [Name / role] |
| Date | 26 June 2026 |
| Status | Draft for internal audit review — not a certified compliance report |
| Distribution | Internal audit / compliance stakeholders |

# 1. Purpose & Scope

This document summarizes the Candour product concept, system architecture, and known compliance, safety, and data-protection considerations, to support internal audit and pre-certification review.

This revision incorporates a significant architecture change — the addition of outward-facing stereo cameras alongside the original inward gaze-tracking concept — which introduces new regulatory and privacy considerations beyond the original gaze-only device. This remains a planning and review document, not a certified compliance report. Formal regulatory testing, legal review, and sign-off by accredited test houses and qualified counsel are required before market release.

# 2. Product Description & Intended Use

Candour is a camera-equipped smart-eyewear bridge module: outward-facing stereo RGB cameras combined with an inward-facing binocular infrared gaze-tracking array, fused on-device to determine what the wearer is looking at and how far away it is. A companion mainboard, carried or placed nearby, receives the processed result.

Components:

- Bridge module — front: stereo cameras, IMU, on-board vision processor; rear: binocular IR gaze array.

- Mainboard companion device — worn or carried separately from the glasses.

- Charging case accessory — fast and ambient charging for the bridge module.

The addition of outward-facing cameras narrows what a “non-medical consumer accessory” framing can cover on its own, and introduces regulatory categories (recording devices, imaging-equipped wearables) that did not apply to the original gaze-only concept. Re-confirm the intended-use statement with this in mind before relying on the rest of this section.

# 3. System Overview

| Component | Role | Notes |
| --- | --- | --- |
| Bridge module — front | Captures stereo imagery and builds a 3D map of the scene | Outward-facing; new in this revision |
| Bridge module — rear | Tracks both eyes’ gaze via reflected infrared, no inward cameras | Inward-facing |
| Mainboard | Receives the fused gaze/scene result, processes and stores it | Carried by the user or left stationary nearby |
| Charging case | Houses charging electronics for the bridge module | Accessory item, not part of the active data path |

Design intent: the bridge module fuses gaze vectors with the 3D scene map on-device, and the wireless link is intended to carry only this fused, compact result — not raw camera frames and not raw photodiode traces.

This is a stated design intent, not yet a verified implementation. It should be confirmed in firmware/code review before being relied on as a privacy mitigation in Sections 5 and 6.

# 4. Regulatory & Standards Considerations

The table below identifies frameworks that are likely relevant given the product description. It is a starting point for due diligence, not a determination of applicability or compliance.

| Area | Likely applicable framework | Why it matters | Status |
| --- | --- | --- | --- |
| Radio emissions | UK: Radio Equipment Regs / UKCA; EU: RED / CE; US: FCC Part 15 | The 2.4 GHz radio is an intentional radiator and typically needs type approval | Not yet tested |
| RF exposure near the body | ICNIRP guidelines / SAR-style assessment | Device is worn on the head; near-body exposure limits can apply even at low transmit power | Not yet assessed |
| Battery safety | IEC 62133-2 (cell/pack safety); UN 38.3 (transport) | Fast-charge battery worn close to the face; charge profile is being re-derived (see engineering spec) | Not yet tested |
| Optical / IR eye safety | IEC 62471; ISO 15004-2 (ophthalmic instruments) | Two to four 940 nm IR LEDs per eye sit close to and continuously illuminate the eye | Not yet assessed — new in this revision |
| Recording-device / surveillance rules | Varies by jurisdiction — wiretap/consent law, bystander-recording rules | Outward-facing cameras can capture identifiable third parties without their knowledge | Not yet assessed — new in this revision |
| Electrical / thermal safety | IEC 62368-1 (or IEC 60950-1 legacy) | General consumer product electrical safety | Not yet assessed |
| Materials | RoHS; REACH | Standard requirement for electronics sold in the UK/EU | Not yet assessed |
| End-of-life | WEEE | Producer take-back and labelling obligations for electronic accessories | Not yet assessed |
| Data protection | UK GDPR / EU GDPR | Now covers both gaze biometric data and bystander camera imagery | See Section 5 |

Engage an accredited test lab and qualified legal counsel to confirm which obligations actually apply, particularly for the two rows newly introduced by the camera addition, before sale.

# 5. Data Privacy & Biometric / Bystander Data Considerations

- Gaze-tracking data can reveal sensitive information about a person and may be treated as biometric data under data protection law if used to identify an individual — carried over from the original gaze-only design.

- New in this revision: the outward stereo cameras capture imagery of the wearer’s environment, which will often include identifiable bystanders. Capturing other people’s images without their knowledge is a separate, often more demanding compliance question than gaze data alone, and is the kind of issue that has drawn regulatory and public scrutiny for other camera-equipped smart glasses products.

- Mitigating design intent: on-device fusion means only a processed gaze/scene result is meant to leave the bridge module, not raw camera frames or raw photodiode traces — a meaningful privacy-by-design choice, but one that needs implementation verification (see Section 3 caveat), and it does not by itself remove the need for a recording indicator or a DPIA.

- Recommended: a visible, active recording indicator (the bridge module’s front status LED is a candidate) so bystanders can tell when the outward cameras are active; a DPIA covering both gaze data and bystander imagery; and a documented data-retention and use policy.

- Encryption: the link uses AES-128. Key-provisioning method (on-chip TRNG-generated session keys vs. a static shared key) still needs to be finalized and verified in code review.

# 6. Safety Risk Register

| Risk | Current design mitigation | Residual risk / open action |
| --- | --- | --- |
| Battery thermal event during fast charge | LTO-class chemistry intended to tolerate high charge rates | Charge profile and capacity are being re-derived for the new power draw; re-test once finalized |
| RF exposure near eyes/face | Default low transmit power when the mainboard is nearby; power only increases when needed | Formal RF exposure assessment not yet performed |
| Continuous near-IR exposure to the eye | 940 nm wavelength and low-power emitters, in line with common commercial eye-tracker practice | New risk in this revision — requires IEC 62471 / ISO 15004-2 testing; emitters sit closer to, and aim more continuously at, the eye than typical AR-glasses use cases, so low apparent power should not be assumed sufficient |
| Unresolved power/thermal budget for the camera + vision-processor stack | None yet — flagged as the top engineering priority | New risk in this revision — battery-safety and thermal conclusions elsewhere in this register cannot be finalized until this is resolved |
| Bystander recording without consent | On-device fusion (raw frames not transmitted); proposed status indicator LED | New risk in this revision — needs an explicit recording-indicator policy and legal review per target jurisdiction |
| Mechanical detachment from glasses | CNC-machined housing direction adopted; attachment method not yet finalized | Define attachment method and run mechanical fatigue testing |
| Small-parts / ingress hazard | Sealed enclosure intended | Confirm enclosure design and ingress protection rating |

# 7. Security Considerations

- Wireless link is encrypted using AES-128.

- No firmware-hardcoded keys is a stated design goal; needs implementation and code review to confirm before audit sign-off.

- Firmware update path (over-the-air or wired) is not yet defined. A signed-firmware update mechanism is recommended for a body-worn, camera-equipped device.

- If any compressed scene imagery or model intermediate output is ever logged or transmitted beyond the fused gaze/object result, that needs separate documentation and review, since it changes the data-protection analysis in Section 5.

# 8. Quality & Reliability Notes

Headline claims that need formal validation before any audit sign-off:

- 48-hour runtime per charge — currently unconfirmed; the underlying power budget is being re-derived to account for the new camera and vision-processor stack (see engineering specification).

- ≤10-minute full recharge — likewise unconfirmed pending the same rework.

Recommended test plan: battery cycle-life testing once the cell is finalized, real-world RF range testing across representative home layouts, near-IR eye-safety testing for the gaze array, and everyday-wear environmental testing (moisture/sweat exposure, drop testing for the bridge module).

# 9. Open Items Before Audit Sign-Off

- Formally define the intended-use statement — now more consequential given the addition of outward-facing cameras.

- Complete a DPIA covering both gaze biometric data and bystander camera capture.

- Re-derive and validate the power/thermal budget for the new front-facing camera and vision-processor stack.

- Commission IEC 62471 / ISO 15004-2 near-infrared eye-safety testing for the gaze-array LEDs.

- Commission RF emissions/exposure and battery safety testing once a working prototype exists.

- Decide and document a bystander-notification approach (e.g. a recording indicator) and review it against applicable recording/consent law per target market.

- Define and implement the firmware key-management and update-signing approach.

# 10. Revision History

| Version | Date | Author | Change description |
| --- | --- | --- | --- |
| v0.1 | 19 June 2026 | [Name] | Initial draft for internal audit review (as “ReArm Snap’n’Track”) |
| v0.2 | 26 June 2026 | [Name] | Renamed product to Candour; architecture updated to add outward-facing stereo cameras and a binocular flat-IR gaze array; added camera-related privacy/regulatory considerations and a near-IR eye-safety risk |

# 11. Approval

Prepared by:  ___________________________	Date:  ___________________

Reviewed by (technical):  ___________________________	Date:  ___________________

Reviewed by (compliance/legal):  ___________________________	Date:  ___________________

Approved by:  ___________________________	Date:  ___________________